SKYE INSTRUMENTS LIMITED

21 Ddole Enterprise Park, Llandrindod Wells, Powys, LD1 6DF UK

Tel: +44 (0) 1597 824811   Fax: +44 (0) 1597 824812

Email: skyemail@skyeinstruments.com     Web:  http://www.skyeinstruments.com

# SKYE TEMPERATURE PROBES

# USING WITH VOLTAGE POWER SUPPLIES OTHER THAN 5V

Skye temperature probes which have a suffix “/R”  have a non linear mV output. They are optimised for use with the Skye DataHog logger or HydroSense meter, which give a regulated 5V excitation supply to the sensor. However, there is a simple ratio calculation when used with other value power supplies. As the sensor’s mV output is directly proportional to the excitation voltage supplied, it is important to provide a stabilised or regulated power supply.

##### CALCULATION

The resistance output is calculated from the temperature variable resistance thermistor T and the 33.2 kohm resistor in parallel as:

Output resistance (R)	= 	( T x 33.2 )  kohm

( T + 33.2 )

Used with an excitation voltage V volts, the output voltage is:

Output voltage (mV) 	=	(V x 1000) mV     x	    R

100 + R

The values given in Table 1 are for use with a 5V excitation supply. From the formula given above it can be seen that the mV output of the probe is directly proportional to the excitation voltage V supplied.

So the given table can be easily converted for other supplies as follows:

For example,

At -20°C with 5V excitation, the probe reads 991.64 mV.

So at 12V excitation, multiply the given mV output by 12 / 5,

the probe will read 2379.94 mV

And at 2V excitation, multiply the given mV output by 2 / 5,

the probe will read 396.66 mV

| TABLE 1 - 10KOHM PRECISION TEMPERATURE THERMISTOR TABLES |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Temp | Thermistor | Output |  | Temp | Thermistor | Output |
|  | value | at 5V |  |  | value | at 5V |
| Deg C | kohms | mV |  | Deg C | kohms | mV |
| -20.00 | 97.08 | 991.64 |  | 41.00 | 5.12 | 212.30 |
| -19.00 | 91.62 | 979.71 |  | 42.00 | 4.92 | 205.37 |
| -18.00 | 86.50 | 967.47 |  | 43.00 | 4.73 | 198.66 |
| -17.00 | 81.70 | 954.91 |  | 44.00 | 4.54 | 192.16 |
| -16.00 | 77.19 | 942.05 |  | 45.00 | 4.37 | 185.88 |
| -15.00 | 72.96 | 928.89 |  | 46.00 | 4.20 | 179.75 |
| -14.00 | 68.98 | 915.46 |  | 47.00 | 4.04 | 173.86 |
| -13.00 | 65.25 | 901.76 |  | 48.00 | 3.89 | 168.22 |
| -12.00 | 61.73 | 887.80 |  | 49.00 | 3.74 | 162.72 |
| -11.00 | 58.43 | 873.61 |  | 50.00 | 3.60 | 157.40 |
| -10.00 | 55.33 | 859.19 |  | 51.00 | 3.47 | 152.21 |
| -9.00 | 52.41 | 844.56 |  | 52.00 | 3.34 | 147.29 |
| -8.00 | 49.66 | 829.75 |  | 53.00 | 3.22 | 142.48 |
| -7.00 | 47.07 | 814.77 |  | 54.00 | 3.10 | 137.84 |
| -6.00 | 44.63 | 799.63 |  | 55.00 | 2.99 | 133.35 |
| -5.00 | 42.33 | 784.36 |  | 56.00 | 2.88 | 129.02 |
| -4.00 | 40.16 | 768.98 |  | 57.00 | 2.77 | 124.83 |
| -3.00 | 38.12 | 753.51 |  | 58.00 | 2.68 | 120.80 |
| -2.00 | 36.19 | 737.96 |  | 59.00 | 2.58 | 116.88 |
| -1.00 | 34.37 | 722.36 |  | 60.00 | 2.49 | 113.12 |
| 0.00 | 32.65 | 706.73 |  | 61.00 | 2.40 | 109.48 |
| 1.00 | 31.03 | 691.13 |  | 62.00 | 2.32 | 105.96 |
| 2.00 | 29.50 | 675.50 |  | 63.00 | 2.24 | 102.56 |
| 3.00 | 28.05 | 659.91 |  | 64.00 | 2.16 | 99.28 |
| 4.00 | 26.68 | 644.37 |  | 65.00 | 2.08 | 96.11 |
| 5.00 | 25.39 | 628.90 |  | 66.00 | 2.01 | 93.05 |
| 6.00 | 24.17 | 613.52 |  | 67.00 | 1.94 | 90.10 |
| 7.00 | 23.01 | 598.25 |  | 68.00 | 1.88 | 87.24 |
| 8.00 | 21.92 | 583.10 |  | 69.00 | 1.81 | 84.49 |
| 9.00 | 20.88 | 568.09 |  | 70.00 | 1.75 | 81.83 |
| 10.00 | 19.90 | 553.24 |  | 71.00 | 1.69 | 79.26 |
| 11.00 | 18.97 | 538.57 |  | 72.00 | 1.64 | 76.78 |
| 12.00 | 18.09 | 524.07 |  | 73.00 | 1.58 | 74.38 |
| 13.00 | 17.25 | 509.77 |  | 74.00 | 1.53 | 72.07 |
| 14.00 | 16.46 | 495.69 |  | 75.00 | 1.48 | 69.82 |
| 15.00 | 15.71 | 481.82 |  | 76.00 | 1.43 | 67.66 |
| 16.00 | 15.00 | 468.18 |  | 77.00 | 1.38 | 65.57 |
| 17.00 | 14.32 | 454.77 |  | 78.00 | 1.34 | 63.57 |
| 18.00 | 13.68 | 441.61 |  | 79.00 | 1.30 | 61.63 |
| 19.00 | 13.07 | 428.71 |  | 80.00 | 1.26 | 59.75 |
| 20.00 | 12.49 | 416.06 |  | 81.00 | 1.22 | 57.93 |
| 21.00 | 11.94 | 403.68 |  | 82.00 | 1.18 | 56.19 |
| 22.00 | 11.42 | 391.56 |  | 83.00 | 1.14 | 54.50 |
| 23.00 | 10.92 | 379.71 |  | 84.00 | 1.10 | 52.86 |
| 24.00 | 10.45 | 368.14 |  | 85.00 | 1.07 | 51.29 |
| 25.00 | 10.00 | 356.84 |  | 86.00 | 1.04 | 49.74 |
| 26.00 | 9.57 | 345.81 |  | 87.00 | 1.00 | 48.27 |
| 27.00 | 9.17 | 335.06 |  | 88.00 | 0.97 | 46.85 |
| 28.00 | 8.78 | 324.58 |  | 89.00 | 0.94 | 45.48 |
| 29.00 | 8.41 | 314.38 |  | 90.00 | 0.92 | 44.14 |
| 30.00 | 8.06 | 304.45 |  | 91.00 | 0.89 | 42.86 |
| 31.00 | 7.72 | 294.79 |  | 92.00 | 0.86 | 41.61 |
| 32.00 | 7.40 | 285.40 |  | 93.00 | 0.84 | 40.41 |
| 33.00 | 7.10 | 276.25 |  | 94.00 | 0.81 | 39.25 |
| 34.00 | 6.81 | 267.38 |  | 95.00 | 0.79 | 38.12 |
| 35.00 | 6.53 | 258.76 |  | 96.00 | 0.76 | 37.03 |
| 36.00 | 6.27 | 250.35 |  | 97.00 | 0.74 | 35.98 |
| 37.00 | 6.02 | 242.33 |  | 98.00 | 0.72 | 34.96 |
| 38.00 | 5.78 | 234.47 |  | 99.00 | 0.70 | 33.98 |
| 39.00 | 5.55 | 226.85 |  | 100.00 | 0.68 | 33.02 |
| 40.00 | 5.33 | 219.46 |  | 101.00 | 0.66 | 32.10 |