SKYE INSTRUMENTS LIMITED

21 Ddole Enterprise Park, Llandrindod Wells, Powys, LD1 6DF UK

Tel: +44 (0) 1597 824811   Fax: +44 (0) 1597 824812

Email: skyemail@skyeinstruments.com     Web:  http://www.skyeinstruments.com

# HOW TO CHECK THE SPECTROSENSE2 METER BATTERY

It is crucial that the SpectroSense2 meter battery is in good condition in order to make accurate measurements, so it is good practice to check the battery voltage before starting.

- Switch on the SpectroSense2 meter, wait 3 seconds until it displays the Main Menu. (If the Main Menu doesn’t appear, simply press the Escape key).

- Use the Down arrow key to choose menu option 4 and press Enter.

- The display will show the voltage of the battery and a “health” message saying whether it is good, medium or low condition.

- Press Escape to return to the Main Menu.

- If the battery voltage is showing low then it requires replacing. Switch off the meter.

- The battery compartment is on the underside of the meter. It takes a PP3 (also called MN1604 or 6LR61) 9V battery. Good quality alkaline batteries are recommended, e.g. Duracell, which will have a lifetime of approximately 40-50 hours of continuous use at 20°C.

- Carefully replace the battery compartment lid ensuring the O ring seal is still in place.