SKYE INSTRUMENTS LIMITED

21 Ddole Enterprise Park, Llandrindod Wells, Powys, LD1 6DF UK

Tel: +44 (0) 1597 824811   Fax: +44 (0) 1597 824812

Email: skyemail@skyeinstruments.com     Web:  http://www.skyeinstruments.com

DATAHOG2 POTENTIOMETER INPUTS

The Skye DataHog2 logger has inputs to accept potentiometer type sensors. The logger supplies a regulated power supply to each sensor.

The number of potentiometer type sensors which can be connected depends on the resistance value of the sensors. The total load on the logger must not exceed an equivalent resistance of 500 ohms, for example:

- Two 1 kohm sensors are permitted, e.g. one Vector W200P windvane plus a 1 kohm evaporation pan gauge

- Ten 5 kohm sensors are permitted.

- Sixteen 10 kohm sensors are permitted, e.g. dendrometers. (Sixteen is the maximum number of single ended voltage hardware channels available on the DataHog2).

For other combinations of different value potentiometers, please check with Skye.

The regulated power supply to the sensors is nominally 1.3V. Each logger is individually calibrated and the exact value is given on the Calibration Certificate.

NOTE – The DataHog2 can only record one windvane sensor in units of degrees. The raw output of all sensors can be datalogged, but the logger configuration cannot automatically convert more than software channel to wind degrees. The sensor’s raw output values can be converted to wind direction degrees in Excel or similar after data download.

Other sensors with a linear Ax+B scaling are not affected, each sensor can be logged in engineering units as usual.