SKYE INSTRUMENTS LIMITED

21 Ddole Enterprise Park, Llandrindod Wells, Powys, LD1 6DF UK

Tel: +44 (0) 1597 824811   Fax: +44 (0) 1597 824812

Email: skyemail@skyeinstruments.com     Web:  http://www.skyeinstruments.com

DATAHOG2

Using with the Delta T SPN1 Sunshine Sensors

The DataHog is compatible with the Delta T SPN1 Sunshine Sensor, but requires a specially configured input socket to record the 3 available outputs. This socket is configured as:

- 0-2.5V for total solar radiation

- 0-2.5V for diffuse solar radiation

- Digital switch for sunshine status (for sunshine hours) where a value of 0 = no sunshine and 1 = sunshine

- /F power supply to the sensor

The DataHog can be configured to record solar radiation in watts m-2 . The sensor is configured ready for use by Skye when supplied as a system with the DataHog

The SPN1 sensor is powered via the DataHog logger, which itself must be connected to an external supply, such as the Skye Mains Hog, Solar Hog or a 12V vehicle battery. The sunshine sensor is also fitted with an internal heater requiring a power supply of 12-15 VDC @ 1.5 A. The DataHog logger is not able to supply this level of power and so the heater supply must be arranged separately.

The total and diffuse solar radiation outputs are usually recorded as average measurements, e.g. the DataHog2’s default setting is to record the average of 30 second “samples” at 30 minutes intervals.

The sunshine status can be recorded in several ways, according to user preference. The status is shown as a zero value for no sunshine and a value of 1 for sunshine.

- An averaged reading of sunshine status, as described above for solar radiation intensity, will give an indication of the percentage of sunshine during the 30 minute period (e.g. a reading of 0.75 indicates the sun was shining for 75% of the time).

- A “spot” recording will give the actual sunshine status at the time of measurement. For this the DataHog2’s sample and logging intervals should be identical, e.g. 30 minute sample and 30 minute logging. Please note that the sunshine status between the 30 minute measurement will not be monitored in this case.

- Log only full sunshine measurements. By using the DataHog2’s “Log Above Threshold” function, it can be set to record only the sunshine status of 1 and ignore the zero readings (e.g. only log values above 0.9). A “spot” logging interval more frequent than 30 minutes is usually chosen as many zero measurement ill be ignored. Please note that choosing non equal logging intervals for different sensor channels can make data difficult to analyse in Excel as data may not always appear in the same columns. However this is not a problem when using Skye’s  SkyeLynx Deluxe software.