SKYE INSTRUMENTS LIMITED

21 Ddole Enterprise Park, Llandrindod Wells, Powys, LD1 6DF UK

Tel: +44 (0) 1597 824811   Fax: +44 (0) 1597 824812

Email: skyemail@skyeinstruments.com     Web:  http://www.skyeinstruments.com

HYDROSENSE METER

Check Calibration Factors

To check the HydroSense meter settings, you will need to have to hand the Calibration Certificates of the sensors which will be connected, and the Calibration and Configuration of the meter itself. Copies are available from Skye if necessary, please let us know the serial numbers of the instruments.

To check the sensor Calibration Factors are entered correctly:

1. Switch on SpectroSense2 meter (a sensor doesn't need to be connected)

2. Press the Next key to display menu 6 Set Parameters, press the 6 key

3. Press the 1 key to choose Browse - Factors

4. If your sensor is entered as Factor 00, press the Next key to display the FS (Full Scale) and OS (Offset). The values should be as per the meter’s Configuration Certificate. Press the Escape key to leave this menu, Escape again returns to the Main Menu.

5. If the meter does not show the values, then you can use the Menu 6 Set Parameters and then 3 Edit Factors to re-enter these values, or you can return the meter to Skye if you prefer.

To check the Differential Amplifier Offset of the meter:

1. From the Main Menu, press the Next key to display menu 6 Set Parameters, press the 6 key

2. Press the Next key to display menu 6 Set Differential Amplifier Offset / Set ID Number, press the 6 key

3. Press the 1 key to choose Edit Differential Amplifier Offset. This display should show the value as per your meter’s Calibration Certificate.

4. Press the Escape key to leave this menu, Escape again returns to the Main Menu.

If the meter values are correct, then please check that they are being used in the readout mode, as follows:

1. From the meter Main Menu, press the 1 key Read / Record

2. Press the Next key to display the current Factors being applied to the readout - this should be 00 for the above example. (The value of the Label is not important as this simple adds a text string to the data in record mode.)

3. If the Factor is not 00, simply change it to 00 (or as appropriate for each sensor) and press enter.