SKYE INSTRUMENTS LIMITED

21 Ddole Enterprise Park, Llandrindod Wells, Powys, LD1 6DF UK

Tel: +44 (0) 1597 824811   Fax: +44 (0) 1597 824812

Email: skyemail@skyeinstruments.com     Web:  http://www.skyeinstruments.com

LIGHT SENSORS

Is my Light Sensor working correctly?

The only accurate method of checking the correct operation of a light sensor, is to compare it’s output against an identical calibrated sensor. However, this is not always easy or convenient, so the guidelines below offer a quick rough check method.

Each Skye light sensor is supplied with a Calibration Certificate, showing the sensor output in current (A or mA) or voltage (mV or V) per unit of light (e.g. A per mol m-2 s-1). The current or voltage output from the sensor can be measured using a multimeter.

The correct configuration of scaling factors in a Skye DataHog logger or SpectroSense meter can also be checked using these guidelines. The following table shows the approximate values of light sensor units to expect from each type of sensor:

## SENSOR TYPE	TYPICAL READINGS ON	ON A CLEAR SUMMER DAY

PAR Special		2,000 mol.m-2.s-1

PAR Quantum		2,000 mol.m-2.s-1

PAR Energy		500 watts.m-2

Lux			100,000 lux

Pyranometer		1,000 watts.m-2

Red 660 nm		150 mol.m-2.s-1

Far-red 730 nm		150 mol.m-2.s-1

Red / Far-red ratio	1.0 around solar noon

UVA			50 watts.m-2

UVB			2 watts.m-2

UVI			>4 UVI or >0.1 watts.m-2

## SENSOR TYPE	VERY APPROXIMATE READINGS

## 30 cm FROM A 100 Watt INCANDESCENT LIGHT BULB

PAR Special		 18 mol.m-2.s-1

PAR Quantum		 19 mol.m-2.s-1

PAR Energy		 4 watts.m-2

Lux			 1000 lux (or 1 klux)

Pyranometer		 30 watts.m-2

Red 660 nm		 2.6 mol.m-2.s-1

Far-red 730 nm		 4.2 mol.m-2.s-1

Red / Far-red ratio	 0.6

## SENSOR TYPE	VERY APPROXIMATE READINGS

## 9 cm FROM A 20 Watt HALOGEN LIGHT BULB

## POWERED BY 12VDC

UVA			 0.8  watts.m-2

UVB			 0.08  watts.m-2

UVI			 0.4 UVI or 0.01 watts.m-2